## 표지(국문)
ISO/IEC 17025:2017 와 KS Q ISO/IEC 17025:2017 에 의한 공인인정범위
삼정산업㈜
경상남도 양산시 안다방길 32-4
전화: 055-384-1611. 팩스: 055-384-1612. 이메일 : twkim@sjcal.kr
교 정
유효기간 만료일 : 2027. 07. 02 | [6 empty] | 인정번호 : KC15-304호
KOLAS 평가결과에 의해 다음의 교정수행에 대하여 공인됩니다.
분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정
|  |  | 40108 | 직류 전원 공급기 | Y
102. 선형치수 |  |  | 40112 | 직류 전압계/차동 전압계 등 | Y | 501. 접촉식온도
10210 | 길이변위계, LVDT | Y |  |  |  | 50101 | 온도 발생장치; 오븐, | Y
10216 | 높이게이지/측정기 | Y | 402. 저항, 용량 및 인덕턴스 |  |  |  | 전기로, 액체항온조, 빙점조,
10234 | 초음파식 두께 측정기 | Y | 40205 | 접지 저항 측정기 | Y |  | 드라이블럭교정기 등
10236 | 피막 두께 측정기 | Y | 40210 | 절연시험기 | Y | 50102 | 온도 지시계; 지시/기록/ | Y
|  |  | 40214 | 저항 측정기, 고저항 측정기 등 | Y |  | 조절계, 온도 교정기 등
106. 기타 길이 관련량 |  |  | 40215 | 저항기, 표준저항, 계단식 저항, | Y | 50103 | 유리제 온도계; 유리제온도계, | N
10601 | 내/외측/기어 이두께 캘리퍼 | Y |  | 고저항 등 |  |  | 백크만 온도계 등
| 캘리퍼 게이지 |  |  |  |  | 50104 | 저항식 온도계; 백금저항온도계, | N
10603 | 실린더/보어 게이지 | Y | 403. 교류 및 교류전력 |  |  |  | 측온저항체, 써미스터 등
10604 | 깊이게이지, 깊이마이크로미터; | Y | 40301 | 교류 전류계 | Y | 50105 | 열팽창식 온도계; 바이메탈 | N
| 다이얼형 포함 |  | 40302 | 클램프형 전류계/전압계 | Y |  | 온도계, 기체 또는
10605 | 다이얼/디지털 게이지 | Y | 40303 | 교류 전압/전류 교정기 | Y |  | 액체 충만식 온도계 등
10609 | 지침 측미기, 테스트 인디케이터 | Y | 40305 | 교류 전류 분류기 | Y | 50106 | 열전대; 귀금속, 비금속, | N
10611 | 3점 마이크로미터 | Y | 40310 | 역률계, 무효율계 등 | Y |  | 순금속, 특수 등
10612 | 내측 마이크로미터 | Y | 40311 | 교류 전력계 | Y
10613 | 외측 마이크로미터 | Y | 40312 | 교류 전원 공급기 | Y | 503. 습도
|  |  | 40313 | 내전압/전기 안전 시험기 | Y | 50302 | 상대습도 습도계; | N
201. 질량 |  |  | 40314 | 전력 기록계 | Y |  | 고분자박막,모발 등
20109 | 전기식 지시저울 | Y | 40318 | 교류 전압계/전위차, 실효치 | Y | 50304 | 온·습도 기록계; | N
20112 | 판수동/플랫트 폼 저울 | Y |  |  |  |  | 자기온습도기록계 등
20113 | 접시 지시저울, 스프링 지시저울 | Y | 404. 기타 직류 및 저주파 측정 |  |  | 50306 | 습도 발생장치; 이압력식/ | Y
20116 | 분동 및 추 | N | 40410 | 전원 주파수계 | Y |  | 이온도식/분류식 습도발생장치,
|  |  | 40411 | 다기능 파형 발생기, 구형파 | Y |  | 항온항습기 등
204. 압력 |  |  |  | 발생기 등
20411 | 게이지압용 압력계; 다이얼, | Y | 40414 | 저주파 임펄스 발생기 | Y
| 디지털, 기록계 등 |  | 40416 | 누설전류 시험기 | Y
20412 | 압력변환기/전송기 | N | 40417 | 직/교류 전자부하 | Y
|  |  | 40419 | 아날로그/디지털 멀티미터; | Y
401. 직류 |  |  |  | 하부 속성 개별 인정
40101 | 직류 전류계 | Y | 40421 | 파형 측정기 | Y
40103 | 직류 전압/전류 교정기 | Y | 40424 | 전압 전류 기록계 | Y
40104 | 전기식온도교정기(센서미포함) | Y | 40425 | 릴레이 시험기 | Y
40105 | 직류용 분류기 | Y | 40426 | LF 신호 발생기 | Y
주석
1. 위 기관은 고정표준실교정 및 현장교정 서비스를 제공한다.
2. 현장교정 서비스를 제공하는 기관은 "KOLAS-SR-007 현장 시험 및 교정 수행을 위한 추가기술요건"을 만족한다.
3. 인정범위 상에 "Y"가 표기된 항목은 현장교정 서비스가 가능하고, "N"이 표기된 항목은 현장교정 서비스가 불가능하다.
4. 측정불확도는 포함인자 k=2를 사용하며, 신뢰수준 약 95 %를 나타내는 확장불확도로 표현되며 정상적인 조건에서 고객에 제공될 수 있는 최소의 측정불확도를 의미한다.
5. 일반적으로, 교정성적서에 기재된 측정불확도는 교정용 표준기, 고객의 장비와 교정환경 등의 영향 때문에, 공인인정범위에 기재된 측정불확도보다 더 크게 표현됨을 유의하여야 한다.

## 교정측정능력(한글)
102. 선형치수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
길이 변위계, LVDT | 10210 |  |  | 게이지 블록
|  | (0 ~ 500) mm |  | / SJ-CAL-10210
높이 게이지/측정기 | 10216 |  |  | 게이지 블록
|  | (0 ~ 1 000) mm |  | / SJ-CAL-10216
초음파식 두께측정기 | 10234 |  |  | 초음파 시편
|  | (0 ~ 200) mm | 8.1 µm | / SJ-CAL-10234
피막 두께측정기 | 10236 |  |  | 두께 측정용 기준시편;코팅형
|  | (0 ~ 1.0) mm | 3.9 µm | / SJ-CAL-10236
106. 기타 길이 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내/외측/기어 이두께 캘리퍼, | 10601 |  |  | 게이지 블록,
캘리퍼게이지 |  |  |  | 캘리퍼 검사기
내/외측 캘리퍼 |  | (0 ~ 1 000) mm |  | / SJ-CAL-10601
게이지 블록
/ SJ-CAL-10601-1
캘리퍼 게이지 외측 |  | (0 ~ 100) mm | 0.82 µm
|  | (100 ~ 300) mm | 7.7 µm
내측 |  | (2.5 ~ 300) mm | 7.7 µm
실린더/보어 게이지 | 10603 |  |  | 다이얼 게이지 시험기
|  | (0 ~ 600) mm | 0.40 µm | / SJ-CAL-10603
깊이 게이지, 깊이 마이크로미터; | 10604 |  |  | 게이지 블록
다이얼형 포함 |  |  |  | / SJ-CAL-10604
깊이 게이지 |  | (0 ~ 50) mm |  | 게이지 블록
/ SJ-CAL-10604-1
(50 ~ 300) mm
마이크로미터 |  | (0 ~ 300) mm
다이얼/디지털 게이지 | 10605 |  |  | 다이얼 게이지 시험기
|  | (0 ~ 50) mm |  | / SJ-CAL-10603
게이지 블록
|  | (50 ~ 100) mm |  | / SJ-CAL-10605
지침 측미기, 테스트 인디케이터 | 10609 |  |  | 다이얼 게이지 시험기
|  | (0 ~ 0.14) mm | 0.59 µm | / SJ-CAL-10609
|  | (0.14 ~ 1.5) mm | 1.5 µm
3점 마이크로미터 | 10611 |  |  | 원통형 링 게이지
|  | (2 ~ 100) mm | 2.4 µm | / SJ-CAL-10611
내측 마이크로미터 | 10612 |  |  | 게이지 블록
봉형 마이크로미터 |  | (100 ~ 1 000) mm |  | / SJ-CAL-10612
게이지 블록
내측 마이크로미터 |  | (5 ~ 300) mm |  | / SJ-CAL-10612-1
106. 기타 길이 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
외측 마이크로미터 | 10613 |  |  | 게이지 블록
|  | (0 ~ 100) mm |  | / SJ-CAL-10613
(100 ~ 1 000) mm
201. 질량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전기식 지시 저울 | 20109 |  |  | 표준 분동
|  | (0 ~ 2) g | 33 μg | / SJ-CAL-20109
|  | (2 ~ 20) g | 61 μg
|  | (20 ~ 50) g | 97 μg
|  | (50 ~ 100) g | 0.13 mg
|  | (100 ~ 200) g | 0.20 mg
|  | (200 ~ 500) g | 0.53 mg
|  | (0.5 ~ 1) kg | 1.3 mg
|  | (1 ~ 2) kg | 2.1 mg
|  | (2 ~ 5) kg | 5.8 mg
|  | (5 ~ 10) kg | 11 mg
|  | (10 ~ 20) kg | 20 mg
|  | (20 ~ 40) kg | 22 mg
|  | (40 ~ 60) kg | 0.12 g
|  | (60 ~ 100) kg | 5.8 g
|  | (100 ~ 500) kg | 32 g
|  | (500 ~ 1 000) kg | 63 g
|  | (1 000 ~ 2 000) kg | 0.13 kg
판수동/플랫트 폼 저울 | 20112 |  |  | 표준 분동
|  | (0 ~ 10) kg | 0.91 g | / SJ-CAL-20112
|  | (10 ~ 20) kg | 1.9 g
|  | (20 ~ 50) kg | 9.5 g
|  | (50 ~ 100) kg | 20 g
|  | (100 ~ 500) kg | 96 g
|  | (500 ~ 1 000) kg | 0.46 kg
|  | (1 000 ~ 2 000) kg | 0.92 kg
접시 지시 저울, | 20113 |  |  | 표준 분동
스프링 지시 저울 등 |  | (0 ~ 1) kg | 0.91 g | / SJ-CAL-20113
|  | (1 ~ 5) kg | 4.6 g
|  | (5 ~ 10) kg | 9.1 g
|  | (10 ~ 50) kg | 46 g
|  | (50 ~ 100) kg | 92 g
201. 질량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
분동 및 추 | 20116 |  |  | 전기식 지시 저울,
|  | (1 mg ~ 20 kg) | M1급 이하 | 표준 분동
|  | 1 mg | 12 μg | / SJ-CAL-20116
|  | 2 mg | 12 μg
|  | 5 mg | 12 μg
|  | 10 mg | 13 μg
|  | 20 mg | 14 μg
|  | 50 mg | 17 μg
|  | 100 mg | 19 μg
|  | 200 mg | 22 μg
|  | 500 mg | 28 μg
|  | 1 g | 35 μg
|  | 2 g | 41 μg
|  | 5 g | 54 μg
|  | 10 g | 61 μg
|  | 20 g | 81 μg
|  | 50 g | 0.10 mg
|  | 100 g | 0.27 mg
|  | 200 g | 0.35 mg
|  | 500 g | 1.2 mg
|  | 1 kg | 1.9 mg
|  | 2 kg | 10 mg
|  | 5 kg | 12 mg
|  | 10 kg | 19 mg
|  | 20 kg | 0.14 g
204. 압력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
게이지압용 압력계; | 20411 |  |  | 분동식압력교정기
다이얼,디지털,기록계 등 |  | (0.02 ~ 7) MPa | 8.6 × 10-5 | / SJ-CAL-20411
|  | (7 ~ 100) MPa | 1.2 × 10-4
|  | (100 ~ 200) MPa | 3.0 × 10-4
압력변환기/전송기 | 20412 |  |  | 분동식압력교정기
|  | (0.02 ~ 7) MPa | 1.6 × 10-4 | / SJ-CAL-20412
|  | (7 ~ 200) MPa | 1.6 × 10-4
401. 직류
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
직류 전류계 | 40101 |  |  | 미터 교정기
직류 전류 |  | (±) |  | /SJ-CAL-40101
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
직류 전압/전류 교정기 | 40103 |  |  | 디지털 멀티미터
직류 전압 |  | (±) |  | /SJ-CAL-40103
|  | 10 mV | 0.69 µV
|  | (10 ~ 100) mV | 1.7 µV
|  | (0.1 ~ 1) V | 6.7 µV
|  | (1 ~ 10) V | 52 µV
|  | (10 ~ 100) V | 0.74 mV
|  | (100 ~ 1 000) V | 9.5 mV
직류 전류 |  | (±)
|  | 10 µA | 7.5 nA
|  | (10 ~ 100) µA | 12 nA
|  | (0.1 ~ 1) mA | 45 nA
|  | (1 ~ 10) mA | 0.50 µA
|  | (10 ~ 100) mA | 9.8 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 4.3 mA
|  | (10 ~ 20) A | 7.6 mA
|  | (20 ~ 30) A | 26 mA
저항 |  | 1 Ω | 26 µΩ
|  | (1 ~ 10) Ω | 0.16 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
|  | (10 ~ 100) kΩ | 1.4 Ω
|  | (0.1 ~ 1) MΩ | 17 Ω
|  | (1 ~ 10) MΩ | 0.35 kΩ
|  | (10 ~ 100) MΩ | 26 kΩ
401. 직류
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전기식온도교정기(센서미포함) | 40104 |  |  | 미터 교정기
온도(Measure) |  |  |  | /SJ-CAL-40104
J-Type |  | (-8.095 ~ 69.553) mV | 1.9 µV
K-Type |  | (-6.458 ~ 54.819) mV | 1.8 µV
T-Type |  | (-6.258 ~ 20.872) mV | 1.4 µV
E-Type |  | (-9.835 ~ 76.373) mV | 1.9 µV
N-Type |  | (-4.345 ~ 47.513) mV | 1.7 µV
R-Type |  | (-0.226 ~ 21.003) mV | 1.4 µV
S-Type |  | (-0.236 ~ 18.609) mV | 1.4 µV
B-Type |  | (0.000 ~ 13.820) mV | 1.4 µV
Pt 100 |  | (18.520 ~ 390.481) Ω | 15 mΩ
온도(Source)
J-Type |  | (-8.095 ~ 69.553) mV | 1.6 µV
K-Type |  | (-6.458 ~ 54.819) mV | 1.6 µV
T-Type |  | (-6.258 ~ 20.872) mV | 1.5 µV
E-Type |  | (-9.835 ~ 76.373) mV | 1.7 µV
N-Type |  | (-4.345 ~ 47.513) mV | 1.5 µV
R-Type |  | (-0.226 ~ 21.003) mV | 1.5 µV
S-Type |  | (-0.236 ~ 18.609) mV | 1.5 µV
B-Type |  | (0.000 ~ 13.820) mV | 1.4 µV
Pt 100 |  | (18.520 ~ 390.481) Ω | 6.8 mΩ
직류용 분류기 | 40105 |  |  | 디지털 멀티미터
저항 |  | 1 mΩ | 0.48 µΩ | /SJ-CAL-40105
|  | (1 ~ 10) mΩ | 3.0 µΩ
|  | (10 ~ 100) mΩ | 16 µΩ
|  | (0.1 ~ 1) Ω | 0.11 mΩ
|  | (1 ~ 10) Ω | 1.1 mΩ
|  | (10 ~ 100) Ω | 12 mΩ
|  | (100 ~ 1 000) Ω | 0.19 Ω
직류 전원 공급기 | 40108 |  |  | 디지털 멀티미터
직류 전압 |  | 100 mV | 6.4 µV | /SJ-CAL-40108
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
직류 전류 |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
401. 직류
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
직류 전압계/차동 전압계 등 | 40112 |  |  | 미터 교정기
직류 전압 |  | (±) |  | /SJ-CAL-40112
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
402. 저항, 용량 및 인덕턴스
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
접지 저항 측정기 | 40205 |  |  | 계단식 저항
교류 전압 |  | 60 Hz |  | /SJ-CAL-40205
|  | 100 mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
교류 전류 |  | 60 Hz
|  | 100 µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
저항 |  | 100 mΩ | 14 µΩ
|  | (0.1 ~ 1) Ω | 0.14 mΩ
|  | (1 ~ 10) Ω | 1.4 mΩ
|  | (10 ~ 100) Ω | 14 mΩ
|  | (0.1 ~ 1) kΩ | 0.14 Ω
|  | (1 ~ 10) kΩ | 1.4 Ω
|  | (10 ~ 100) kΩ | 14 Ω
부하 저항 |  | 100 mΩ | 1.3 mΩ
|  | 300 mΩ | 3.6 mΩ
|  | 500 mΩ | 5.9 mΩ
절연 시험기 | 40210 |  |  | 계단식 저항
직류 전압 |  | 1 V | 62 µV | /SJ-CAL-40210
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
402. 저항, 용량 및 인덕턴스
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
절연 시험기 | 40210 |  |  | 계단식 저항
교류 전압 |  | 60 Hz |  | /SJ-CAL-40210
|  | 1 V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
절연 전압 |  | 1 V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 5) kV | 32 V
|  | (5 ~ 10) kV | 61 V
저항 |  | 1 kΩ | 1.4 Ω
|  | (1 ~ 10) kΩ | 14 Ω
|  | (10 ~ 100) kΩ | 0.14 kΩ
|  | (0.1 ~ 1) MΩ | 1.4 kΩ
|  | (1 ~ 10) MΩ | 14 kΩ
|  | (10 ~ 100) MΩ | 1.2 MΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
|  | (0.1 ~ 1) TΩ | 35 GΩ
시간 |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
|  | (100 ~ 1 000) s | 0.58 s
저항 측정기, 고저항 측정기등 | 40214 |  |  | 표준저항
저항 |  | 1 mΩ | 67 nΩ | /SJ-CAL-40214
|  | (1 ~ 10) mΩ | 0.67 µΩ
|  | (10 ~ 100) mΩ | 6.7 µΩ
|  | (0.1 ~ 1) Ω | 30 µΩ
|  | (1 ~ 10) Ω | 0.14 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
|  | (10 ~ 100) kΩ | 1.4 Ω
|  | (0.1 ~ 1) MΩ | 20 Ω
|  | (1 ~ 10) MΩ | 0.26 kΩ
|  | (10 ~ 100) MΩ | 2.8 kΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
402. 저항, 용량 및 인덕턴스
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
저항기, 표준저항, | 40215 |  |  | 디지털 멀티미터
계단식 저항, 고저항 등 |  |  |  | /SJ-CAL-40215
저항 |  | 1 mΩ | 0.67 µΩ
|  | (1 ~ 10) mΩ | 1.7 µΩ
|  | (10 ~ 100) mΩ | 16 µΩ
|  | (0.1 ~ 1) Ω | 0.16 mΩ
|  | (1 ~ 10) Ω | 1.6 mΩ
|  | (10 ~ 100) Ω | 1.3 mΩ
|  | (0.1 ~ 1) kΩ | 12 mΩ
|  | (1 ~ 10) kΩ | 0.12 Ω
|  | (10 ~ 100) kΩ | 1.3 Ω
|  | (0.1 ~ 1) MΩ | 16 Ω
|  | (1 ~ 10) MΩ | 0.35 kΩ
|  | (10 ~ 100) MΩ | 26 kΩ
|  | (0.1 ~ 1) GΩ | 2.7 MΩ
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전류계 | 40301 |  |  | 미터 교정기
교류 전류 |  | (50 Hz ~ 1 kHz) |  | /SJ-CAL-40301
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
(1 ~ 5) kHz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
클램프형 전류계/전압계 | 40302 |  |  | 미터 교정기
직류 전압 |  | 10 mV | 1.3 µV | /SJ-CAL-40302
|  | (10 ~ 100) mV | 6.5 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
교류 전압 |  | 60 Hz
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
직류 전류 |  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 20 nA
|  | (0.1 ~ 1) mA | 0.13 µA
|  | (1 ~ 10) mA | 1.2 µA
|  | (10 ~ 100) mA | 12 µA
|  | (0.1 ~ 1) A | 0.17 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 0.11 A
|  | (100 ~ 1 000) A | 1.4 A
|  | (1 000 ~ 1 500) A | 2.1 A
교류 전류 |  | 60 Hz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 100) A | 0.40 A
|  | (100 ~ 1 000) A | 1.7 A
|  | (1 000 ~ 1 500) A | 2.4 A
저항 |  | 1 Ω | 68 µΩ
|  | (1 ~ 10) Ω | 0.63 mΩ
|  | (10 ~ 100) Ω | 6.3 mΩ
|  | (0.1 ~ 1) kΩ | 63 mΩ
|  | (1 ~ 10) kΩ | 0.63 Ω
|  | (10 ~ 100) kΩ | 6.3 Ω
|  | (0.1 ~ 1) MΩ | 64 Ω
|  | (1 ~ 10) MΩ | 0.67 kΩ
|  | (10 ~ 100) MΩ | 6.7 kΩ
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전압/전류 교정기 | 40303 |  |  | 디지털 멀티미터
교류 전압 |  | (40 Hz ~ 1 kHz) |  | /SJ-CAL-40303
|  | 10 mV | 7.4 µV
|  | (10 ~ 100) mV | 19 µV
|  | (0.1 ~ 1) V | 0.12 mV
|  | (1 ~ 10) V | 1.2 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.18 V
(1 ~ 10) kHz
|  | 10 mV | 8.0 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.6 mV
|  | (10 ~ 100) V | 16 mV
교류 전류 |  | (40 Hz ~ 1 kHz)
|  | 10 µA | 27 nA
|  | (10 ~ 100) µA | 45 nA
|  | (0.1 ~ 1) mA | 0.42 µA
|  | (1 ~ 10) mA | 4.2 µA
|  | (10 ~ 100) mA | 42 µA
|  | (0.1 ~ 1) A | 0.57 mA
|  | (1 ~ 10) A | 12 mA
|  | (10 ~ 20) A | 34 mA
|  | (20 ~ 30) A | 0.13 A
(1 ~ 5) kHz
|  | 10 µA | 27 nA
|  | (10 ~ 100) µA | 83 nA
|  | (0.1 ~ 1) mA | 0.77 µA
|  | (1 ~ 10) mA | 7.5 µA
|  | (10 ~ 100) mA | 73 µA
|  | (0.1 ~ 1) A | 0.98 mA
교류 전류 분류기 | 40305 |  |  | 디지털 멀티미터
교류 저항 |  | 60 Hz |  | /SJ-CAL-40305
|  | 1 mΩ | 3.3 µΩ
|  | (1 ~ 10) mΩ | 3.8 µΩ
|  | (10 ~ 100) mΩ | 34 µΩ
|  | (0.1 ~ 1) Ω | 0.28 mΩ
|  | (1 ~ 10) Ω | 3.9 mΩ
|  | (10 ~ 100) Ω | 40 mΩ
|  | (100 ~ 1 000) Ω | 0.56 Ω
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
역률계, 무효율계 등 | 40310 |  |  | 미터 교정기
역률 |  | (60 Hz) |  | /SJ-CAL-40310
|  | 1 | 0.000 7
|  | 0.9(Lead, Lag) | 0.001 1
|  | 0.8(Lead, Lag) | 0.001 4
|  | 0.7(Lead, Lag) | 0.001 6
|  | 0.6(Lead, Lag) | 0.001 8
|  | 0.5(Lead, Lag) | 0.001 9
|  | 0.4(Lead, Lag) | 0.002 0
|  | 0.3(Lead, Lag) | 0.002 1
|  | 0.2(Lead, Lag) | 0.002 1
|  | 0.1(Lead, Lag) | 0.002 1
교류 전력계 | 40311 |  |  | 미터 교정기
직류 전압 |  | 10 mV | 1.3 µV | /SJ-CAL-40311
|  | (10 ~ 100) mV | 6.5 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
교류 전압 |  | 60 Hz
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
직류 전류 |  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 20 nA
|  | (0.1 ~ 1) mA | 0.13 µA
|  | (1 ~ 10) mA | 1.2 µA
|  | (10 ~ 100) mA | 12 µA
|  | (0.1 ~ 1) A | 0.17 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
교류 전류 |  | 60 Hz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전력계 | 40311 |  |  | 미터 교정기
교류 전력 |  | 60 Hz |  | /SJ-CAL-40311
|  | 6 W | 19 mW
|  | (6 ~ 12) W | 19 mW
|  | (12 ~ 18) W | 21 mW
|  | (18 ~ 24) W | 21 mW
|  | (24 ~ 30) W | 22 mW
|  | (30 ~ 36) W | 23 mW
|  | (36 ~ 42) W | 24 mW
|  | (42 ~ 48) W | 25 mW
|  | (48 ~ 54) W | 27 mW
|  | (54 ~ 60) W | 28 mW
|  | (60 ~ 72) W | 37 mW
|  | (72 ~ 84) W | 41 mW
|  | (84 ~ 96) W | 44 mW
|  | (96 ~ 108) W | 47 mW
|  | (108 ~ 120) W | 50 mW
|  | (120 ~ 144) W | 69 mW
|  | (144 ~ 168) W | 74 mW
교류 전력 |  | 60 Hz
|  | (168 ~ 192) W | 79 mW
|  | (192 ~ 216) W | 84 mW
|  | (216 ~ 240) W | 88 mW
|  | (240 ~ 300) W | 0.22 W
|  | (300 ~ 360) W | 0.24 W
|  | (360 ~ 420) W | 0.25 W
|  | (420 ~ 480) W | 0.27 W
|  | (480 ~ 540) W | 0.29 W
|  | (540 ~ 600) W | 0.30 W
|  | (600 ~ 720) W | 0.37 W
|  | (720 ~ 840) W | 0.40 W
|  | (840 ~ 960) W | 0.44 W
|  | (960 ~ 1 080) W | 0.47 W
|  | (1 080 ~ 1 200) W | 0.54 W
|  | (1 200 ~ 1 440) W | 0.71 W
|  | (1 440 ~ 1 680) W | 0.78 W
|  | (1 680 ~ 1 920) W | 0.84 W
|  | (1 920 ~ 2 160) W | 0.90 W
|  | (2 160 ~ 2 400) W | 0.96 W
|  | (2 400 ~ 2 880) W | 3.3 W
|  | (2 880 ~ 3 360) W | 3.8 W
|  | (3 360 ~ 3 840) W | 4.4 W
|  | (3 840 ~ 4 320) W | 4.9 W
|  | (4 320 ~ 4 800) W | 5.4 W
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전력계 | 40311 |  |  | 미터 교정기
역률 |  | 60 Hz |  | /SJ-CAL-40311
|  | 1 | 0.000 7
|  | 0.9(Lead, Lag) | 0.001 1
|  | 0.8(Lead, Lag) | 0.001 4
|  | 0.7(Lead, Lag) | 0.001 6
|  | 0.6(Lead, Lag) | 0.001 8
|  | 0.5(Lead, Lag) | 0.001 9
|  | 0.4(Lead, Lag) | 0.002 0
|  | 0.3(Lead, Lag) | 0.002 1
|  | 0.2(Lead, Lag) | 0.002 1
|  | 0.1(Lead, Lag) | 0.002 1
교류 전원 공급기 | 40312 |  |  | 디지털 멀티미터
교류 전압 |  | (40 Hz ~ 1 kHz) |  | /SJ-CAL-40312
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
교류 전류 |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
주파수 |  | 10 Hz | 0.62 mHz
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
|  | (1 ~ 10) kHz | 0.62 Hz
|  | (10 ~ 100) kHz | 6.2 Hz
|  | (0.1 ~ 1) MHz | 62 Hz
내전압/전기 안전 시험기 | 40313 |  |  | 고전압 미터
직류 전압 |  | 0.1 kV | 3.2 V | /SJ-CAL-40313
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내전압/전기 안전 시험기 | 40313 |  |  | 고전압 미터
교류 전압 |  | (50 ~ 60) Hz |  | /SJ-CAL-40313
|  | 0.1 kV | 3.2 V
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
직류 차단 전류 |  | 0.5 mA | 4.3 µA
|  | (0.5 ~ 1) mA | 7.1 µA
|  | (1 ~ 2) mA | 13 µA
|  | (2 ~ 5) mA | 31 µA
|  | (5 ~ 10) mA | 71 µA
|  | (10 ~ 20) mA | 0.13 mA
|  | (20 ~ 50) mA | 0.31 mA
|  | (50 ~ 100) mA | 0.71 mA
교류 차단 전류 |  | (50 ~ 60) Hz
|  | 0.5 mA | 4.3 µA
|  | (0.5 ~ 1) mA | 7.1 µA
|  | (1 ~ 2) mA | 13 µA
|  | (2 ~ 5) mA | 31 µA
|  | (5 ~ 10) mA | 71 µA
|  | (10 ~ 20) mA | 0.13 mA
|  | (20 ~ 50) mA | 0.31 mA
|  | (50 ~ 100) mA | 0.71 mA
절연 전압 |  | 0.1 kV | 3.2 V
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내전압/전기 안전 시험기 | 40313 |  |  | 고전압 미터
저항 |  | 1 kΩ | 1.4 Ω | /SJ-CAL-40313
|  | (1 ~ 10) kΩ | 14 Ω
|  | (10 ~ 100) kΩ | 0.14 kΩ
|  | (0.1 ~ 1) MΩ | 1.4 kΩ
|  | (1 ~ 10) MΩ | 14 kΩ
|  | (10 ~ 100) MΩ | 1.2 MΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
|  | (0.1 ~ 1) TΩ | 35 GΩ
부하 저항 |  | 100 mΩ | 1.3 mΩ
|  | 300 mΩ | 3.6 mΩ
|  | 500 mΩ | 5.9 mΩ
시간 |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
|  | (100 ~ 1 000) s | 0.58 s
전력 기록계 | 40314 |  |  | 미터 교정기
교류 전력 |  | 60 Hz |  | /SJ-CAL-40314
|  | 6 W | 19 mW
|  | (6 ~ 12) W | 19 mW
|  | (12 ~ 18) W | 21 mW
|  | (18 ~ 24) W | 21 mW
|  | (24 ~ 30) W | 22 mW
|  | (30 ~ 36) W | 23 mW
|  | (36 ~ 42) W | 24 mW
|  | (42 ~ 48) W | 25 mW
|  | (48 ~ 54) W | 27 mW
|  | (54 ~ 60) W | 28 mW
|  | (60 ~ 72) W | 37 mW
|  | (72 ~ 84) W | 41 mW
|  | (84 ~ 96) W | 44 mW
|  | (96 ~ 108) W | 47 mW
|  | (108 ~ 120) W | 50 mW
|  | (120 ~ 144) W | 69 mW
|  | (144 ~ 168) W | 74 mW
|  | (168 ~ 192) W | 79 mW
|  | (192 ~ 216) W | 84 mW
|  | (216 ~ 240) W | 88 mW
|  | (240 ~ 300) W | 0.22 W
|  | (300 ~ 360) W | 0.24 W
|  | (360 ~ 420) W | 0.25 W
|  | (420 ~ 480) W | 0.27 W
|  | (480 ~ 540) W | 0.29 W
|  | (540 ~ 600) W | 0.30 W
403. 교류 및 교류 전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전력 기록계 | 40314 |  |  | 미터 교정기
교류 전력 |  | 60 Hz |  | /SJ-CAL-40314
|  | (600 ~ 720) W | 0.37 W
|  | (720 ~ 840) W | 0.40 W
|  | (840 ~ 960) W | 0.44 W
|  | (960 ~ 1 080) W | 0.47 W
|  | (1 080 ~ 1 200) W | 0.54 W
|  | (1 200 ~ 1 440) W | 0.71 W
|  | (1 440 ~ 1 680) W | 0.78 W
|  | (1 680 ~ 1 920) W | 0.84 W
|  | (1 920 ~ 2 160) W | 0.90 W
|  | (2 160 ~ 2 400) W | 0.96 W
|  | (2 400 ~ 2 880) W | 3.3 W
|  | (2 880 ~ 3 360) W | 3.8 W
|  | (3 360 ~ 3 840) W | 4.4 W
|  | (3 840 ~ 4 320) W | 4.9 W
|  | (4 320 ~ 4 800) W | 5.4 W
교류 전압계 | 40318 |  |  | 미터 교정기
교류 전압 |  | (50 Hz ~ 10 kHz) |  | /SJ-CAL-40318
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전원 주파수계 | 40410 |  |  | 미터 교정기
주파수 |  | 10 Hz | 0.24 mHz | /SJ-CAL-40410
|  | (10 ~ 100) Hz | 2.4 mHz
|  | (0.1 ~ 1) kHz | 24 mHz
|  | (1 ~ 10) kHz | 0.24 Hz
|  | (10 ~ 100) kHz | 2.4 Hz
|  | (0.1 ~ 1) MHz | 24 Hz
다기형 파형 발생기, | 40411 |  |  | 주파수 계수기
구형파 발생기 등 |  |  |  | /SJ-CAL-40411
주파수 |  | 10 Hz | 6.5 µHz
|  | (10 ~ 100) Hz | 65 µHz
|  | (0.1 ~ 1) kHz | 0.65 mHz
|  | (1 ~ 10) kHz | 6.5 mHz
|  | (10 ~ 100) kHz | 65 mHz
|  | (0.1 ~ 1) MHz | 0.65 Hz
|  | (1 ~ 10) MHz | 6.5 Hz
|  | (10 ~ 100) MHz | 65 Hz
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
다기형 파형 발생기, | 40411 |  |  | 주파수 계수기
구형파 발생기 등 |  |  |  | /SJ-CAL-40411
교류 전압 |  | (40 Hz ~ 1 kHz)
(Sine, Square, Ramp) |  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.14 mV
|  | (1 ~ 10) V | 1.3 mV
출력 레벨 |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
출력 평탄성 |  | 1 V
|  | 10 Hz | 0.34 mV
|  | (10 ~ 100) Hz | 0.13 mV
|  | (0.1 ~ 1) kHz | 0.13 mV
|  | (1 ~ 10) kHz | 0.17 mV
|  | (10 ~ 100) kHz | 0.69 mV
직류 옵셋 |  | (±)
|  | 100 mV | 6.4 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
상승/하강 시간 |  | 100 ns | 0.12 ns
|  | (0.1 ~ 1) µs | 1.2 ns
|  | (1 ~ 10) µs | 12 ns
|  | (10 ~ 100) µs | 0.12 µs
|  | (0.1 ~ 1) ms | 1.2 µs
저주파 임펄스 발생기 | 40414 |  |  | 고압 프로브
펄스 전압 |  | 0.1 kV | 14 V | /SJ-CAL-40414
|  | (0.1 ~ 0.5) kV | 28 V
|  | (0.5 ~ 1) kV | 52 V
|  | (1 ~ 2) kV | 0.11 kV
|  | (2 ~ 3) kV | 0.17 kV
|  | (3 ~ 4) kV | 0.19 kV
|  | (4 ~ 5) kV | 0.26 kV
|  | (5 ~ 10) kV | 0.52 kV
|  | (10 ~ 15) kV | 0.85 kV
|  | (15 ~ 20) kV | 1.1 kV
누설전류 시험기 | 40416 |  |  | 미터 교정기
직류 전압 |  | 10 mV | 6.3 µV | /SJ-CAL-40416
|  | (10 ~ 100) mV | 62 µV
|  | (0.1 ~ 1) V | 0.62 mV
|  | (1 ~ 10) V | 6.2 mV
|  | (10 ~ 100) V | 62 mV
|  | (100 ~ 1 000) V | 0.62 V
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
누설전류 시험기 | 40416 |  |  | 미터 교정기
/SJ-CAL-40416
교류 전압 |  | (50 Hz ~ 1 kHz)
|  | 10 mV | 12 µV
|  | (10 ~ 100) mV | 65 µV
|  | (0.1 ~ 1) V | 0.63 mV
|  | (1 ~ 10) V | 6.3 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
직류 전류 |  | 10 µA | 11 nA
|  | (10 ~ 100) µA | 64 nA
|  | (0.1 ~ 1) mA | 0.63 µA
|  | (1 ~ 10) mA | 6.2 µA
|  | (10 ~ 100) mA | 62 µA
교류 전류 |  | (50 Hz ~ 1 kHz)
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 81 nA
|  | (0.1 ~ 1) mA | 0.71 µA
|  | (1 ~ 10) mA | 7.0 µA
|  | (10 ~ 100) mA | 65 µA
직/교류 전자부하 | 40417 |  |  | 디지털 멀티미터
직류 전압 |  | 100 mV | 6.4 µV | /SJ-CAL-40417
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
직류 전류 |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
교류 전압 |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
교류 전류 |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
아날로그/디지털 멀티미터; | 40419 |  |  | 미터 교정기
하부 속성 개별인정 |  |  |  | /SJ-CAL-40419
직류 전압 |  | (±)
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
교류 전압 |  | (50 Hz ~ 10 kHz)
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
직류 전류 |  | (±)
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
교류 전류 |  | (50 Hz ~ 1 kHz)
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
(1 ~ 5) kHz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
아날로그/디지털 멀티미터; | 40419 |  |  | 미터 교정기
하부 속성 개별인정 |  |  |  | /SJ-CAL-40419
저항 |  | 1 Ω | 30 µΩ
|  | (1 ~ 10) Ω | 0.14 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
|  | (10 ~ 100) kΩ | 1.4 Ω
|  | (0.1 ~ 1) MΩ | 20 Ω
|  | (1 ~ 10) MΩ | 0.26 kΩ
|  | (10 ~ 100) MΩ | 2.8 kΩ
파형 측정기 | 40421 |  |  | 미터 교정기
직류 전압 |  | (±) |  | /SJ-CAL-40421
|  | 1 mV | 47 µV
|  | (1 ~ 10) mV | 51 µV
|  | (10 ~ 100) mV | 93 µV
|  | (0.1 ~ 1) V | 0.52 mV
|  | (1 ~ 10) V | 4.7 mV
|  | (10 ~ 100) V | 47 mV
|  | (100 ~ 120) V | 56 mV
구형파 전압 |  | 1 kHz
|  | 2 mV | 49 µV
|  | (2 ~ 6) mV | 52 µV
|  | (6 ~ 12) mV | 58 µV
|  | (12 ~ 30) mV | 75 µV
|  | (30 ~ 60) mV | 0.11 mV
|  | (60 ~ 120) mV | 0.17 mV
|  | (120 ~ 300) mV | 0.34 mV
|  | (300 ~ 600) mV | 0.61 mV
|  | (0.6 ~ 1.2) V | 1.2 mV
|  | (1.2 ~ 3) V | 2.9 mV
|  | (3 ~ 6) V | 5.7 mV
|  | (6 ~ 12) V | 12 mV
|  | (12 ~ 30) V | 28 mV
|  | (30 ~ 60) V | 56 mV
|  | (60 ~ 120) V | 0.12 V
대역폭 |  | 600 mV
|  | 50 kHz | 8.5 mV
|  | (50 ~ 100) kHz | 9.0 mV
|  | (0.1 ~ 10) MHz | 14 mV
|  | (10 ~ 60) MHz | 20 mV
|  | (60 ~ 600) MHz | 21 mV
|  | (600 ~ 1 000) MHz | 26 mV
|  | (1 000 ~ 1 100) MHz | 27 mV
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
파형 측정기 | 40421 |  |  | 미터 교정기
시간 |  | 1 ns | 0.62 ps | /SJ-CAL-40421
|  | (1 ~ 5) ns | 0.62 ps
|  | (5 ~ 50) ns | 6.2 ps
|  | (50 ~ 500) ns | 62 ps
|  | (0.5 ~ 5) µs | 0.62 ns
|  | (5 ~ 50) µs | 6.2 ns
|  | (50 ~ 500) µs | 62 ns
|  | (0.5 ~ 5) ms | 0.62 µs
|  | (5 ~ 50) ms | 6.2 µs
|  | (50 ~ 500) ms | 62 µs
|  | (0.5 ~ 5) s | 0.62 ms
전압 전류 기록계 | 40424 |  |  | 미터 교정기
직류 전압 |  | (±) |  | /SJ-CAL-40424
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
교류 전압 |  | (50 Hz ~ 10 kHz)
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
직류 전류 |  | (±)
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전압 전류 기록계 | 40424 |  |  | 미터 교정기
교류 전류 |  | (50 Hz ~ 1 kHz) |  | /SJ-CAL-40424
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
(1 ~ 5) kHz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
릴레이 시험기 | 40425 |  |  | 디지털 멀티미터
직류 전압 |  | 100 mV | 6.4 µV | /SJ-CAL-40425
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
직류 전류 |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
교류 전압 |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
교류 전류 |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
404. 기타 직류 및 저주파 측정
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
릴레이 시험기 | 40425 |  |  | 디지털 멀티미터
주파수 |  | 10 Hz | 0.62 mHz | /SJ-CAL-40425
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
시간 |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
LF 신호 발생기 | 40426 |  |  | 주파수 계수기
주파수 |  | 10 Hz | 0.62 mHz | /SJ-CAL-40426
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
|  | (1 ~ 10) kHz | 0.62 Hz
|  | (10 ~ 100) kHz | 6.2 Hz
|  | (0.1 ~ 1) MHz | 62 Hz
출력 평탄성 |  | 1 V
|  | 10 Hz | 0.70 mV
|  | (10 ~ 100) Hz | 0.63 mV
|  | (0.1 ~ 1) kHz | 0.63 mV
|  | (1 ~ 10) kHz | 0.64 mV
|  | (10 ~ 100) kHz | 0.92 mV
감쇠량 |  | 1 kHz
0 dBV |  | 1 V | 0.63 mV
-10 dBV |  | 316.2 mV | 91 µV
-20 dBV |  | 100 mV | 64 µV
501. 접촉식 온도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
온도 발생장치; 오븐, 전기로, | 50101 |  |  | SPRT, Thermocouple
액체항온조, 빙점조, 드라이블록 |  |  |  | / SJ-CAL-50101
교정기 등
온도조절형 챔버 및 오븐 |  | (-39.5 ~ 250) ℃ | 1.4 ℃
드라이블록 교정기 |  | (-39.5 ~ 400) ℃ | 0.06 ℃
온도지시계; 지시/기록/조절계, | 50102 |  |  | SPRT, Thermocouple,
온도교정기 등 |  |  |  | Decade Resistor,
온도지시/기록/조절계 |  |  |  | Calibrator
(센서포함) |  | (-39.5 ~ 250) ℃ | 0.04 ℃ | / SJ-CAL-50102
|  | (250 ~ 1 100) ℃ | 0.9 ℃
|  | (1 100 ~ 1 300) ℃ | 2.6 ℃
(센서제외) |  | (-39.5 ~ 1 300) ℃ | 0.1 ℃
유리제 온도계; 유리제 온도계, | 50103 |  |  | SPRT
벡크만 온도계 등 |  | (-39.5 ~ 250) ℃ | 0.05 ℃ | / SJ-CAL-50103
501. 접촉식 온도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
저항식 온도계; 백금저항온도계, | 50104 |  |  | SPRT
측온저항체, 써미스터 등 |  |  |  | / SJ-CAL-50104
측온저항체 |  | (-39.5 ~ 250) ℃ | 0.03 ℃
열 팽창식 온도계; 바이메탈온도계, | 50105 |  |  | SPRT
기체 또는 액체충만온도계 등 |  |  |  | / SJ-CAL-50105
바이메탈온도계 |  | (-39.5 ~ 250) ℃ | 0.30 ℃
열전대; 귀금속, 비금속, 순금속, | 50106 |  |  | SPRT, Thermocouple
특수 등 |  |  |  | / SJ-CAL-50106
비금속, 귀금속 |  | (-39.5 ~ 250) ℃ | 0.3 ℃
|  | (250 ~ 1 100) ℃ | 1.3 ℃
|  | (1 100 ~ 1 300) ℃ | 2.7 ℃
503. 습도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
상대습도 습도계; 고분자박막, | 50302 |  |  | 노점 습도계/
모발 등 |  |  |  | SJ-CAL-50302
고분자박막(상대습도) |  | (30 ~ 95) % R.H. | 2.2
고분자박막(온도) |  | (-30 ~ 70) ℃ | 0.6
모발(상대습도) |  | (30 ~ 80) % R.H. | 4.2
(온도) |  | (-20 ~ 50) ℃ | 1.4
온·습도 기록계; 자기온습도 | 50304 |  |  | 노점 습도계/
기록계 등 |  |  |  | SJ-CAL-50304
(상대습도) |  | (30 ~ 80) % R.H. | 4.2
(온도) |  | (-20 ~ 50) ℃ | 1.4
습도 발생장치; 이압력식/이온 | 50306 |  |  | 노점습도계/
도식/분류식 습도발생장치, |  |  |  | SJ-CAL-50306
항온항습기 등
항온항습기(상대습도) |  | (30 ~ 60) % R.H. | 2.4
|  | (60 ~ 97) % R.H. | 3.4
(온도) |  | (-39.5 ~ 250) ℃ | 0.7
√("0." 〖"3" 4〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"2.51" 〗^2+("0.003 " 0×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"15" 〗^2+("0.003 0" ×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"0.1" 5〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"1." 8〗^2+("0.003 1" ×𝑙〖0)〗^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"11.3" 〗^2+("0.004 " 0×𝑙〖0)〗^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"0.77" 〗^2+("0.003 " 2×𝑙0)^2 ) "μm "
(𝑙0 단위 :" mm")
√(〖"0.71" 〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"1.8" 〗^2+("0.003 0" ×𝑙〖0)〗^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"1." 5〗^2+("0.003 0" ×𝑙0)^2 ) " μm "
(𝑙0 단위 :" mm")
√(〖"0.92" 〗^2+("0.002 9" ×𝑙0)^2 ) "μm "
(𝑙0 단위 :" mm")
√(〖"1." 9〗^2+("0.003 0" ×𝑙0)^2 ) "μm "
(𝑙0 단위 :" mm")

## 표지(영문)
SCOPE OF ACCREDITATION TO ISO/IEC 17025:2017 & KS Q ISO/IEC 17025:2017
Sam Jeong industry Co., Ltd.
32-4, Andabang-gil, Yangsan-si, Gyeongsangnam-do, Republic of Korea
Phone No : 82-55-384-1611, FAX : 82-55-384-1612, E-mail : twkim@sjcal.kr
CALIBRATION
Valid To : July. 02, 2027 | [5 empty] | Accreditation No : KC15-304
In recognition of the successful completion of the KOLAS evaluation process,
accreditation is granted to this laboratory to perform the following calibrations
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
|  |  | 40103 | DC voltage/current calibrators | Y | 40424 | Volt/Current recorders | Y
102. Linear dimension |  |  | 40104 | Electrical temperature | Y | 40425 | Relay test sets | Y
10210 | Extensometer, linear | Y |  | calibrators |  | 40426 | LF signal generators | Y
| displacement transducers |  | 40105 | DC current shunts | Y
10216 | Height gauges/ | Y | 40108 | DC power supplies | Y | 501. Contact thermometry
| measuring machines |  | 40112 | DC voltmeters | Y | 50101 | Temperature generators: oven | Y
10234 | Ultrasonic thickness gauges | Y | 402. Resistance, Capacitance and |  |  |  | furnaces,isothermal liquid baths,
10236 | Coating thickness testers | Y | Inductance |  |  |  | ice-point baths,
|  |  | 40205 | Earth testers | Y |  | dry-block calibrators
106. Various dimensional |  |  | 40210 | Insulation testers | Y | 50102 | Temperature indicators/recorders/ | Y
10601 | Inside/Outside/Gear tooth | Y | 40214 | Resistance meters | Y |  | controllers, temperature
| calipers, Caliper gauges |  | 40215 | Resistors | Y |  | calibrators
10603 | Cylinder/Bore gauges | Y
10604 | Depth gauges, | Y | 403. AC voltage, current & power |  |  | 50103 | Glass thermometers:liquid-in-glass, | N
| Depth micrometers |  | 40301 | AC ammeters | Y |  | Beckmann
10605 | Dial/Digital gauges | Y | 40302 | Clamp ammeters/voltmeters | Y | 50104 | Resistance thermometers: SPRT, | N
10609 | Micro indicators, | Y | 40303 | AC voltage/current | Y |  | IPRT,thermistors, etc
| Test indicators | Y |  | calibrators |  | 50105 | Thermal expansion thermometers: | N
10611 | 3-point micrometers | Y | 40305 | AC current shunts | Y |  | bimeter, gas or liquid type
10612 | Inside micrometers | Y | 40310 | Power factor meters | Y | 50106 | Thermocouples: noble metal, | N
10613 | Outside micrometers | Y | 40311 | AC power meters | Y |  | base metal,pure metal,
|  |  | 40312 | AC power supplies | Y |  | special type, etc.
201. Mass |  |  | 40313 | Puncture/safety testers | Y
20109 | Electric balances | Y | 40314 | Power recorders | Y | 503. Humidity
20112 | Plarform scale balances | Y | 40318 | AC voltmeters | Y | 50302 | Relative humidity | N
20113 | Spring scale balances | Y |  |  |  |  | hygrometers; polimer
20116 | Weights | N | 404. Other DC & LF Measurements |  |  |  | thinfilm, hair, etc.
|  |  | 40410 | Line frequency meter | Y | 50304 | Temperature humidity | N
204. Pressure |  |  | 40411 | Function generators | Y |  | recorders ;
20411 | Gague pressure gauges | Y | 40414 | LF impulse generators | Y |  | Hygrothermograph, etc
20412 | Pressure transducers / | N | 40416 | Leakage current testers | Y | 50306 | Humidity generators; | Y
| transmitters |  | 40417 | Electronic AC/DC loads | Y |  | two-pressure, two-temperature,
|  |  | 40419 | Analogue/Digital | Y |  | flow mixing humidity gererator,
401. DC voltage & current |  |  |  | multimeters |  |  | constant temperature and
40101 | DC ammeters | Y | 40421 | Oscilloscopes | Y |  | humidity chamber, etc.
Note
1. This laboratory provides calibration services in permanent standard laboratory and at on-site.
2. Laboratory conducts on-site calibration should meet requirements of KOLAS-SR-007.
3. On-site calibration is allowed to items with marking 'Y', not allowed to items with marking 'N'.
4. Measurement uncertainty normally is quoted as an expanded uncertainty at a coverage probability of 95 %, which usually requires the use of a coverage factor of k=2. It expresses the lowest uncertainty of measurement that can be provided by accredited calbration laboratories in normal conditions.
5.Due to the calibration environment such as reference standards or customers’ facilities, it is note that uncertainty of measurement on a calibration certificate may be expressed larger than measurement uncertainty on scope of accreditation in general.

## 교정측정능력(영문)
102. Linear dimension
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Extensometers, linear | 10210 |  |  | Gauge blocks
displacement transducers |  | (0 ~ 500) mm |  | /SJ-CAL-10210
Height gauges/ | 10216 |  |  | Gauge Blocks
measuring machines |  | (0 ~ 1 000) mm |  | /SJ-CAL-10216
Ultrasonic thickness gauges | 10234 |  |  | Ultrasonic specimens
|  | (0 ~ 200) mm | 8.1 µm | /SJ-CAL-10234
Coating thickness testers | 10236 |  |  | coating thickness specimens
|  | (0 ~ 1.0) mm | 3.9 µm | /SJ-CAL-10236
106.Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Inside/Outside/Gear tooth | 10601 |  |  | Gauge Blocks,
calipers, Caliper gauges |  |  |  | Caliper Checker
Inside/Outside caliper |  | (0 ~ 1 000) mm |  | /SJ-CAL-10601
caliper gauges |  |  |  | Gauge Blocks
Outside |  | (0 ~ 100) mm | 0.82 µm | /SJ-CAL-10601-1
|  | (100 ~ 300) mm | 7.7 µm
Inside |  | (2.5 ~ 300) mm | 7.7 µm
Cylinder/Bore gauges | 10603 |  |  | Dial gauge tester
|  | (0 ~ 600) mm | 0.40 µm | /SJ-CAL-10603
Depth gauges, | 10604
Depth micrometers
Depth gauges |  | (0 ~ 50) mm |  | Gauge Blocks
/SJ-CAL-10604
(50 ~ 300) mm
Depth micrometers |  | (0 ~ 300) mm |  | Gauge Blocks
/SJ-CAL-10604-1
Dial/Digital gauges | 10605 |  |  | Dial gauge tester,
|  | (0 ~ 50) mm |  | Gauge Blocks
/SJ-CAL-10603
|  | (50 ~ 100) mm |  | / SJ-CAL-10605
Micro indicators,Test indicators | 10609
|  | (0 ~ 0.14) mm | 0.59 µm | Dial gauge tester
|  | (0.14 ~ 1.5) mm | 1.5 µm | /SJ-CAL-10609
3-points, Micrometers | 10611 |  |  | Cylindrical ring gauge
|  | (2 ~ 100) mm | 2.4 µm | /SJ-CAL-10611
106.Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Inside micrometers | 10612
bar type Micrometers |  | (100 ~ 1 000) mm |  | Gauge Blocks
/SJ-CAL-10612
Inside micrometers |  | (5 ~ 300) mm |  | Gauge Blocks
/SJ-CAL-10612-1
Outside micrometers | 10613 |  |  | Gauge Blocks
|  | (0 ~ 100) mm |  | /SJ-CAL-10613
(100 ~ 1 000) mm
201. Mass
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electric balances | 20109 |  |  | Weights
|  | (0 ~ 2) g | 33 μg | /SJ-CAL-20109
|  | (2 ~ 20) g | 61 μg
|  | (20 ~ 50) g | 97 μg
|  | (50 ~ 100) g | 0.13 mg
|  | (100 ~ 200) g | 0.20 mg
|  | (200 ~ 500) g | 0.53 mg
|  | (0.5 ~ 1) kg | 1.3 mg
|  | (1 ~ 2) kg | 2.1 mg
|  | (2 ~ 5) kg | 5.8 mg
|  | (5 ~ 10) kg | 11 mg
|  | (10 ~ 20) kg | 20 mg
|  | (20 ~ 40) kg | 22 mg
|  | (40 ~ 60) kg | 0.12 g
|  | (60 ~ 100) kg | 5.8 g
|  | (100 ~ 500) kg | 32 g
|  | (500 ~ 1 000) kg | 63 g
|  | (1 000 ~ 2 000) kg | 0.13 kg
Plarform scale balances | 20112 |  |  | Weights
|  | (0 ~ 10) kg | 0.91 g | /SJ-CAL-20112
|  | (10 ~ 20) kg | 1.9 g
|  | (20 ~ 50) kg | 9.5 g
|  | (50 ~ 100) kg | 20 g
|  | (100 ~ 500) kg | 96 g
|  | (500 ~ 1 000) kg | 0.46 kg
|  | (1 000 ~ 2 000) kg | 0.92 kg
Spring scale balances | 20113 |  |  | Weights
|  | (0 ~ 1) kg | 0.91 g | /SJ-CAL-20113
|  | (1 ~ 5) kg | 4.6 g
|  | (5 ~ 10) kg | 9.1 g
|  | (10 ~ 50) kg | 46 g
|  | (50 ~ 100) kg | 92 g
201. Mass
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Weights | 20116 |  |  | Weights, Electric
|  | (1 mg ~ 20 kg) | Class M1 or less | balances
|  | 1 mg | 12 μg | /SJ-CAL-20116
|  | 2 mg | 12 μg
|  | 5 mg | 12 μg
|  | 10 mg | 13 μg
|  | 20 mg | 14 μg
|  | 50 mg | 17 μg
|  | 100 mg | 19 μg
|  | 200 mg | 22 μg
|  | 500 mg | 28 μg
|  | 1 g | 35 μg
|  | 2 g | 41 μg
|  | 5 g | 54 μg
|  | 10 g | 61 μg
|  | 20 g | 81 μg
|  | 50 g | 0.10 mg
|  | 100 g | 0.27 mg
|  | 200 g | 0.35 mg
|  | 500 g | 1.2 mg
|  | 1 kg | 1.9 mg
|  | 2 kg | 10 mg
|  | 5 kg | 12 mg
|  | 10 kg | 19 mg
|  | 20 kg | 0.14 g
204. Pressure
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Gauge pressure gauges | 20411 |  |  | Deadweight Tester /SJ-CAL-20411
|  | (0.02 ~ 7) MPa | 8.6 × 10-5
|  | (7 ~ 100) MPa | 1.2 × 10-4
|  | (100 ~ 200) MPa | 3.0 × 10-4
Pressure transducers / | 20412 |  |  | Deadweight Tester /SJ-CAL-20412
transmitters |  | (0.02 ~ 7) MPa | 1.6 × 10-4
|  | (7 ~ 200) MPa | 1.6 × 10-4
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC ammeters | 40101 |  |  | Calibrator
DC Current |  | (±) |  | /SJ-CAL-40101
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
DC voltage/current calibrators | 40103 |  |  | Digital Multimeter
DC Voltage |  | (±) |  | /SJ-CAL-40103
|  | 10 mV | 0.69 µV
|  | (10 ~ 100) mV | 1.7 µV
|  | (0.1 ~ 1) V | 6.7 µV
|  | (1 ~ 10) V | 52 µV
|  | (10 ~ 100) V | 0.74 mV
|  | (100 ~ 1 000) V | 9.5 mV
DC Current |  | (±)
|  | 10 µA | 7.5 nA
|  | (10 ~ 100) µA | 12 nA
|  | (0.1 ~ 1) mA | 45 nA
|  | (1 ~ 10) mA | 0.50 µA
|  | (10 ~ 100) mA | 9.8 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 4.3 mA
|  | (10 ~ 20) A | 7.6 mA
|  | (20 ~ 30) A | 26 mA
Resistance |  | 1 Ω | 26 µΩ
|  | (1 ~ 10) Ω | 0.16 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
|  | (10 ~ 100) kΩ | 1.4 Ω
|  | (0.1 ~ 1) MΩ | 17 Ω
|  | (1 ~ 10) MΩ | 0.35 kΩ
|  | (10 ~ 100) MΩ | 26 kΩ
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electrical temperature calibrators | 40104 |  |  | Calibrator
Temperature(Measure) |  |  |  | /SJ-CAL-40104
J-Type |  | (-8.095 ~ 69.553) mV | 1.9 µV
K-Type |  | (-6.458 ~ 54.819) mV | 1.8 µV
T-Type |  | (-6.258 ~ 20.872) mV | 1.4 µV
E-Type |  | (-9.835 ~ 76.373) mV | 1.9 µV
N-Type |  | (-4.345 ~ 47.513) mV | 1.7 µV
R-Type |  | (-0.226 ~ 21.003) mV | 1.4 µV
S-Type |  | (-0.236 ~ 18.609) mV | 1.4 µV
B-Type |  | (0.000 ~ 13.820) mV | 1.4 µV
Pt 100 |  | (18.520 ~ 390.481) Ω | 15 mΩ
Temperature(Source)
J-Type |  | (-8.095 ~ 69.553) mV | 1.6 µV
K-Type |  | (-6.458 ~ 54.819) mV | 1.6 µV
T-Type |  | (-6.258 ~ 20.872) mV | 1.5 µV
E-Type |  | (-9.835 ~ 76.373) mV | 1.7 µV
N-Type |  | (-4.345 ~ 47.513) mV | 1.5 µV
R-Type |  | (-0.226 ~ 21.003) mV | 1.5 µV
S-Type |  | (-0.236 ~ 18.609) mV | 1.5 µV
B-Type |  | (0.000 ~ 13.820) mV | 1.4 µV
Pt 100 |  | (18.520 ~ 390.481) Ω | 6.8 mΩ
DC current shunts | 40105 |  |  | Digital Multimeter
Resistance |  | 1 mΩ | 0.48 µΩ | /SJ-CAL-40105
|  | (1 ~ 10) mΩ | 3.0 µΩ
|  | (10 ~ 100) mΩ | 16 µΩ
|  | (0.1 ~ 1) Ω | 0.11 mΩ
|  | (1 ~ 10) Ω | 1.1 mΩ
|  | (10 ~ 100) Ω | 12 mΩ
|  | (100 ~ 1 000) Ω | 0.19 Ω
DC power supplies | 40108 |  |  | Digital Multimeter
DC Voltage |  | 100 mV | 6.4 µV | /SJ-CAL-40108
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
DC Current |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
401. DC voltage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC voltmeters | 40112 |  |  | Calibrator
DC Voltage |  | (±) |  | /SJ-CAL-40112
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
402. Resistance, capacitance and inductance
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Earth testers | 40205 |  |  | Decade Resistor
AC Voltage |  | 60 Hz |  | /SJ-CAL-40205
|  | 100 mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
AC Current |  | 60 Hz
|  | 100 µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
Resistance |  | 100 mΩ | 14 µΩ
|  | (0.1 ~ 1) Ω | 0.14 mΩ
|  | (1 ~ 10) Ω | 1.4 mΩ
|  | (10 ~ 100) Ω | 14 mΩ
|  | (0.1 ~ 1) kΩ | 0.14 Ω
|  | (1 ~ 10) kΩ | 1.4 Ω
|  | (10 ~ 100) kΩ | 14 Ω
Load Resistance |  | 100 mΩ | 1.3 mΩ
|  | 300 mΩ | 3.6 mΩ
|  | 500 mΩ | 5.9 mΩ
402. Resistance, capacitance and inductance
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Insulation testers | 40210 |  |  | Decade Resistor
DC Voltage |  | 1 V | 62 µV | /SJ-CAL-40210
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
AC Voltage |  | 60 Hz
|  | 1 V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
Insulation Voltage |  | 1 V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 5) kV | 32 V
|  | (5 ~ 10) kV | 61 V
Resistance |  | 1 kΩ | 1.4 Ω
|  | (1 ~ 10) kΩ | 14 Ω
|  | (10 ~ 100) kΩ | 0.14 kΩ
|  | (0.1 ~ 1) MΩ | 1.4 kΩ
|  | (1 ~ 10) MΩ | 14 kΩ
|  | (10 ~ 100) MΩ | 1.2 MΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
|  | (0.1 ~ 1) TΩ | 35 GΩ
Time |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
|  | (100 ~ 1 000) s | 0.58 s
Resistance meters | 40214 |  |  | Standard Resistance
Resistance |  | 1 mΩ | 67 nΩ | /SJ-CAL-40214
|  | (1 ~ 10) mΩ | 0.67 µΩ
|  | (10 ~ 100) mΩ | 6.7 µΩ
|  | (0.1 ~ 1) Ω | 30 µΩ
|  | (1 ~ 10) Ω | 0.14 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
402. Resistance, capacitance and inductance
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Resistance meters | 40214 |  |  | Standard Resistance
Resistance |  | (10 ~ 100) kΩ | 1.4 Ω | /SJ-CAL-40214
|  | (0.1 ~ 1) MΩ | 20 Ω
|  | (1 ~ 10) MΩ | 0.26 kΩ
|  | (10 ~ 100) MΩ | 2.8 kΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
Resistors | 40215 |  |  | Digital Multimeter
Resistance |  | 1 mΩ | 0.67 µΩ | /SJ-CAL-40215
|  | (1 ~ 10) mΩ | 1.7 µΩ
|  | (10 ~ 100) mΩ | 16 µΩ
|  | (0.1 ~ 1) Ω | 0.16 mΩ
|  | (1 ~ 10) Ω | 1.6 mΩ
|  | (10 ~ 100) Ω | 1.3 mΩ
|  | (0.1 ~ 1) kΩ | 12 mΩ
|  | (1 ~ 10) kΩ | 0.12 Ω
|  | (10 ~ 100) kΩ | 1.3 Ω
|  | (0.1 ~ 1) MΩ | 16 Ω
|  | (1 ~ 10) MΩ | 0.35 kΩ
|  | (10 ~ 100) MΩ | 26 kΩ
|  | (0.1 ~ 1) GΩ | 2.7 MΩ
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC ammeters | 40301 |  |  | Calibrator
AC Current |  | (50 Hz ~ 1 kHz) |  | /SJ-CAL-40301
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC ammeters | 40301 |  |  | Calibrator
AC Current |  | (1 ~ 5) kHz |  | /SJ-CAL-40301
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator
DC Voltage |  | 10 mV | 1.3 µV | /SJ-CAL-40302
|  | (10 ~ 100) mV | 6.5 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
AC Voltage |  | 60 Hz
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
DC Current |  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 20 nA
|  | (0.1 ~ 1) mA | 0.13 µA
|  | (1 ~ 10) mA | 1.2 µA
|  | (10 ~ 100) mA | 12 µA
|  | (0.1 ~ 1) A | 0.17 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 0.11 A
|  | (100 ~ 1 000) A | 1.4 A
|  | (1 000 ~ 1 500) A | 2.1 A
AC Current |  | 60 Hz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 100) A | 0.40 A
|  | (100 ~ 1 000) A | 1.7 A
|  | (1 000 ~ 1 500) A | 2.4 A
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Clamp ammeters/voltmeters | 40302 |  |  | Calibrator
Resistance |  | 1 Ω | 68 µΩ | /SJ-CAL-40302
|  | (1 ~ 10) Ω | 0.63 mΩ
|  | (10 ~ 100) Ω | 6.3 mΩ
|  | (0.1 ~ 1) kΩ | 63 mΩ
|  | (1 ~ 10) kΩ | 0.63 Ω
|  | (10 ~ 100) kΩ | 6.3 Ω
|  | (0.1 ~ 1) MΩ | 64 Ω
|  | (1 ~ 10) MΩ | 0.67 kΩ
|  | (10 ~ 100) MΩ | 6.7 kΩ
AC voltage/current calibrators | 40303 |  |  | Digital Multimeter
AC Voltage |  | (40 Hz ~ 1 kHz) |  | /SJ-CAL-40303
|  | 10 mV | 7.4 µV
|  | (10 ~ 100) mV | 19 µV
|  | (0.1 ~ 1) V | 0.12 mV
|  | (1 ~ 10) V | 1.2 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.18 V
(1 ~ 10) kHz
|  | 10 mV | 8.0 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.6 mV
|  | (10 ~ 100) V | 16 mV
AC Current |  | (40 Hz ~ 1 kHz)
|  | 10 µA | 27 nA
|  | (10 ~ 100) µA | 45 nA
|  | (0.1 ~ 1) mA | 0.42 µA
|  | (1 ~ 10) mA | 4.2 µA
|  | (10 ~ 100) mA | 42 µA
|  | (0.1 ~ 1) A | 0.57 mA
|  | (1 ~ 10) A | 12 mA
|  | (10 ~ 20) A | 34 mA
|  | (20 ~ 30) A | 0.13 A
(1 ~ 5) kHz
|  | 10 µA | 27 nA
|  | (10 ~ 100) µA | 83 nA
|  | (0.1 ~ 1) mA | 0.77 µA
|  | (1 ~ 10) mA | 7.5 µA
|  | (10 ~ 100) mA | 73 µA
|  | (0.1 ~ 1) A | 0.98 mA
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC current shunts | 40305 |  |  | Digital Multimeter
AC Resistance |  | 60 Hz |  | /SJ-CAL-40305
|  | 1 mΩ | 3.3 µΩ
|  | (1 ~ 10) mΩ | 3.8 µΩ
|  | (10 ~ 100) mΩ | 34 µΩ
|  | (0.1 ~ 1) Ω | 0.28 mΩ
|  | (1 ~ 10) Ω | 3.9 mΩ
|  | (10 ~ 100) Ω | 40 mΩ
|  | (100 ~ 1 000) Ω | 0.56 Ω
Power factor meters | 40310 |  |  | Calibrator
Power Factor |  | (60 Hz) |  | /SJ-CAL-40310
|  | 1 | 0.000 7
|  | 0.9(Lead, Lag) | 0.001 1
|  | 0.8(Lead, Lag) | 0.001 4
|  | 0.7(Lead, Lag) | 0.001 6
|  | 0.6(Lead, Lag) | 0.001 8
|  | 0.5(Lead, Lag) | 0.001 9
|  | 0.4(Lead, Lag) | 0.002 0
|  | 0.3(Lead, Lag) | 0.002 1
|  | 0.2(Lead, Lag) | 0.002 1
|  | 0.1(Lead, Lag) | 0.002 1
AC power meters | 40311 |  |  | Calibrator
DC Voltage |  | 10 mV | 1.3 µV | /SJ-CAL-40311
|  | (10 ~ 100) mV | 6.5 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.3 mV
|  | (100 ~ 1 000) V | 63 mV
AC Voltage |  | 60 Hz
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 22 µV
|  | (0.1 ~ 1) V | 0.16 mV
|  | (1 ~ 10) V | 1.5 mV
|  | (10 ~ 100) V | 16 mV
|  | (100 ~ 1 000) V | 0.24 V
DC Current |  | 10 µA | 8.2 nA | /SJ-CAL-40311
|  | (10 ~ 100) µA | 20 nA
|  | (0.1 ~ 1) mA | 0.13 µA
|  | (1 ~ 10) mA | 1.2 µA
|  | (10 ~ 100) mA | 12 µA
|  | (0.1 ~ 1) A | 0.17 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311 |  |  | Calibrator
AC Current |  | 60 Hz |  | /SJ-CAL-40311
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.5 µA
|  | (10 ~ 100) mA | 22 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.4 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
AC Wattage |  | 60 Hz
|  | 6 W | 19 mW
|  | (6 ~ 12) W | 19 mW
|  | (12 ~ 18) W | 21 mW
|  | (18 ~ 24) W | 21 mW
|  | (24 ~ 30) W | 22 mW
|  | (30 ~ 36) W | 23 mW
|  | (36 ~ 42) W | 24 mW
|  | (42 ~ 48) W | 25 mW
|  | (48 ~ 54) W | 27 mW
|  | (54 ~ 60) W | 28 mW
|  | (60 ~ 72) W | 37 mW
|  | (72 ~ 84) W | 41 mW
|  | (84 ~ 96) W | 44 mW
|  | (96 ~ 108) W | 47 mW
|  | (108 ~ 120) W | 50 mW
|  | (120 ~ 144) W | 69 mW
|  | (144 ~ 168) W | 74 mW
AC Wattage |  | 60 Hz
|  | (168 ~ 192) W | 79 mW
|  | (192 ~ 216) W | 84 mW
|  | (216 ~ 240) W | 88 mW
|  | (240 ~ 300) W | 0.22 W
|  | (300 ~ 360) W | 0.24 W
|  | (360 ~ 420) W | 0.25 W
|  | (420 ~ 480) W | 0.27 W
|  | (480 ~ 540) W | 0.29 W
|  | (540 ~ 600) W | 0.30 W
|  | (600 ~ 720) W | 0.37 W
|  | (720 ~ 840) W | 0.40 W
|  | (840 ~ 960) W | 0.44 W
|  | (960 ~ 1 080) W | 0.47 W
|  | (1 080 ~ 1 200) W | 0.54 W
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311 |  |  | Calibrator
AC Wattage |  | 60 Hz |  | /SJ-CAL-40311
|  | (1 200 ~ 1 440) W | 0.71 W
|  | (1 440 ~ 1 680) W | 0.78 W
|  | (1 680 ~ 1 920) W | 0.84 W
|  | (1 920 ~ 2 160) W | 0.90 W
|  | (2 160 ~ 2 400) W | 0.96 W
|  | (2 400 ~ 2 880) W | 3.3 W
|  | (2 880 ~ 3 360) W | 3.8 W
|  | (3 360 ~ 3 840) W | 4.4 W
|  | (3 840 ~ 4 320) W | 4.9 W
|  | (4 320 ~ 4 800) W | 5.4 W
Power Factor |  | 60 Hz
|  | 1 | 0.000 7
|  | 0.9(Lead, Lag) | 0.001 1
|  | 0.8(Lead, Lag) | 0.001 4
|  | 0.7(Lead, Lag) | 0.001 6
|  | 0.6(Lead, Lag) | 0.001 8
|  | 0.5(Lead, Lag) | 0.001 9
|  | 0.4(Lead, Lag) | 0.002 0
|  | 0.3(Lead, Lag) | 0.002 1
|  | 0.2(Lead, Lag) | 0.002 1
|  | 0.1(Lead, Lag) | 0.002 1
AC power supplies | 40312 |  |  | Digital Multimeter
AC Voltage |  | (40 Hz ~ 1 kHz) |  | /SJ-CAL-40312
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
AC Current |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
Frequency |  | 10 Hz | 0.62 mHz
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
|  | (1 ~ 10) kHz | 0.62 Hz
|  | (10 ~ 100) kHz | 6.2 Hz
|  | (0.1 ~ 1) MHz | 62 Hz
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Puncture/safety testers | 40313 |  |  | High Voltage Meter
DC Voltage |  | 0.1 kV | 3.2 V | /SJ-CAL-40313
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
AC Voltage |  | (50 ~ 60) Hz
|  | 0.1 kV | 3.2 V
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
DC Breaking Current |  | 0.5 mA | 4.3 µA
|  | (0.5 ~ 1) mA | 7.1 µA
|  | (1 ~ 2) mA | 13 µA
|  | (2 ~ 5) mA | 31 µA
|  | (5 ~ 10) mA | 71 µA
|  | (10 ~ 20) mA | 0.13 mA
|  | (20 ~ 50) mA | 0.31 mA
|  | (50 ~ 100) mA | 0.71 mA
AC Breaking Current |  | (50 ~ 60) Hz
|  | 0.5 mA | 4.3 µA
|  | (0.5 ~ 1) mA | 7.1 µA
|  | (1 ~ 2) mA | 13 µA
|  | (2 ~ 5) mA | 31 µA
|  | (5 ~ 10) mA | 71 µA
|  | (10 ~20) mA | 0.13 mA
|  | (20 ~ 50) mA | 0.31 mA
|  | (50 ~ 100) mA | 0.71 mA
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Puncture/safety testers | 40313 |  |  | High Voltage Meter
Insulation Voltage |  | 0.1 kV | 3.2 V | /SJ-CAL-40313
|  | (0.1 ~ 0.5) kV | 5.4 V
|  | (0.5 ~ 1) kV | 8.2 V
|  | (1 ~ 2) kV | 14 V
|  | (2 ~ 3) kV | 20 V
|  | (3 ~ 4) kV | 26 V
|  | (4 ~ 5) kV | 32 V
|  | (5 ~ 6) kV | 38 V
|  | (6 ~ 7) kV | 43 V
|  | (7 ~ 8) kV | 49 V
|  | (8 ~ 9) kV | 55 V
|  | (9 ~ 10) kV | 61 V
Resistance |  | 1 kΩ | 1.4 Ω
|  | (1 ~ 10) kΩ | 14 Ω
|  | (10 ~ 100) kΩ | 0.14 kΩ
|  | (0.1 ~ 1) MΩ | 1.4 kΩ
|  | (1 ~ 10) MΩ | 14 kΩ
|  | (10 ~ 100) MΩ | 1.2 MΩ
|  | (0.1 ~ 1) GΩ | 12 MΩ
|  | (1 ~ 10) GΩ | 0.12 GΩ
|  | (10 ~ 100) GΩ | 1.2 GΩ
|  | (0.1 ~ 1) TΩ | 35 GΩ
Load Resistance |  | 100 mΩ | 1.3 mΩ
|  | 300 mΩ | 3.6 mΩ
|  | 500 mΩ | 5.9 mΩ
Time |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
|  | (100 ~ 1 000) s | 0.58 s
Power recorders | 40314 |  |  | Calibrator
AC Wattage |  | 60 Hz |  | /SJ-CAL-40314
|  | 6 W | 19 mW
|  | (6 ~ 12) W | 19 mW
|  | (12 ~ 18) W | 21 mW
|  | (18 ~ 24) W | 21 mW
|  | (24 ~ 30) W | 22 mW
|  | (30 ~ 36) W | 23 mW
|  | (36 ~ 42) W | 24 mW
|  | (42 ~ 48) W | 25 mW
|  | (48 ~ 54) W | 27 mW
|  | (54 ~ 60) W | 28 mW
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Power recorders | 40314 |  |  | Calibrator
AC Wattage |  | 60 Hz |  | /SJ-CAL-40314
|  | (60 ~ 72) W | 37 mW
|  | (72 ~ 84) W | 41 mW
|  | (84 ~ 96) W | 44 mW
|  | (96 ~ 108) W | 47 mW
|  | (108 ~ 120) W | 50 mW
|  | (120 ~ 144) W | 69 mW
|  | (144 ~ 168) W | 74 mW
|  | (168 ~ 192) W | 79 mW
|  | (192 ~ 216) W | 84 mW
|  | (216 ~ 240) W | 88 mW
|  | (240 ~ 300) W | 0.22 W
|  | (300 ~ 360) W | 0.24 W
|  | (360 ~ 420) W | 0.25 W
|  | (420 ~ 480) W | 0.27 W
|  | (480 ~ 540) W | 0.29 W
|  | (540 ~ 600) W | 0.30 W
|  | (600 ~ 720) W | 0.37 W
|  | (720 ~ 840) W | 0.40 W
|  | (840 ~ 960) W | 0.44 W
|  | (960 ~ 1 080) W | 0.47 W
|  | (1 080 ~ 1 200) W | 0.54 W
|  | (1 200 ~ 1 440) W | 0.71 W
|  | (1 440 ~ 1 680) W | 0.78 W
|  | (1 680 ~ 1 920) W | 0.84 W
|  | (1 920 ~ 2 160) W | 0.90 W
|  | (2 160 ~ 2 400) W | 0.96 W
|  | (2 400 ~ 2 880) W | 3.3 W
|  | (2 880 ~ 3 360) W | 3.8 W
|  | (3 360 ~ 3 840) W | 4.4 W
|  | (3 840 ~ 4 320) W | 4.9 W
|  | (4 320 ~ 4 800) W | 5.4 W
AC voltmeters | 40318 |  |  | Calibrator
AC Voltage |  | (50 Hz ~ 10 kHz) |  | /SJ-CAL-40318
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Line frequency meters | 40410 |  |  | Calibrator
Frequency |  | 10 Hz | 0.24 mHz | /SJ-CAL-40410
|  | (10 ~ 100) Hz | 2.4 mHz
|  | (0.1 ~ 1) kHz | 24 mHz
|  | (1 ~ 10) kHz | 0.24 Hz
|  | (10 ~ 100) kHz | 2.4 Hz
|  | (0.1 ~ 1) MHz | 24 Hz
Function generators | 40411 |  |  | Frequency Counter
Frequency |  | 10 Hz | 6.5 µHz | /SJ-CAL-40411
|  | (10 ~ 100) Hz | 65 µHz
|  | (0.1 ~ 1) kHz | 0.65 mHz
|  | (1 ~ 10) kHz | 6.5 mHz
|  | (10 ~ 100) kHz | 65 mHz
|  | (0.1 ~ 1) MHz | 0.65 Hz
|  | (1 ~ 10) MHz | 6.5 Hz
|  | (10 ~ 100) MHz | 65 Hz
AC Voltage |  | (40 Hz ~ 1 kHz)
(Sine, Square, Ramp) |  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.14 mV
|  | (1 ~ 10) V | 1.3 mV
Output Level |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
Level Flatness |  | 1 V
|  | 10 Hz | 0.34 mV
|  | (10 ~ 100) Hz | 0.13 mV
|  | (0.1 ~ 1) kHz | 0.13 mV
|  | (1 ~ 10) kHz | 0.17 mV
|  | (10 ~ 100) kHz | 0.69 mV
DC Offset |  | (±)
|  | 100 mV | 6.4 µV
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
Rise/Fall Time |  | 100 ns | 0.12 ns
|  | (0.1 ~ 1) µs | 1.2 ns
|  | (1 ~ 10) µs | 12 ns
|  | (10 ~ 100) µs | 0.12 µs
|  | (0.1 ~ 1) ms | 1.2 µs
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF impulse generators | 40414 |  |  | High Voltage Probe
Pulse Voltage |  | 0.1 kV | 14 V | /SJ-CAL-40414
|  | (0.1 ~ 0.5) kV | 28 V
|  | (0.5 ~ 1) kV | 52 V
|  | (1 ~ 2) kV | 0.11 kV
|  | (2 ~ 3) kV | 0.17 kV
|  | (3 ~ 4) kV | 0.19 kV
|  | (4 ~ 5) kV | 0.26 kV
|  | (5 ~ 10) kV | 0.52 kV
|  | (10 ~ 15) kV | 0.85 kV
|  | (15 ~ 20) kV | 1.1 kV
Leakage current testers | 40416 |  |  | Calibrator
DC Voltage |  | 10 mV | 6.3 µV | /SJ-CAL-40416
|  | (10 ~ 100) mV | 62 µV
|  | (0.1 ~ 1) V | 0.62 mV
|  | (1 ~ 10) V | 6.2 mV
|  | (10 ~ 100) V | 62 mV
|  | (100 ~ 1 000) V | 0.62 V
AC Voltage |  | (50 Hz ~ 1 kHz)
|  | 10 mV | 12 µV
|  | (10 ~ 100) mV | 65 µV
|  | (0.1 ~ 1) V | 0.63 mV
|  | (1 ~ 10) V | 6.3 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
DC Current |  | 10 µA | 11 nA
|  | (10 ~ 100) µA | 64 nA
|  | (0.1 ~ 1) mA | 0.63 µA
|  | (1 ~ 10) mA | 6.2 µA
|  | (10 ~ 100) mA | 62 µA
AC Current |  | (50 Hz ~ 1 kHz)
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 81 nA
|  | (0.1 ~ 1) mA | 0.71 µA
|  | (1 ~ 10) mA | 7.0 µA
|  | (10 ~ 100) mA | 65 µA
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electronic AC/DC loads | 40417 |  |  | Digital Multimeter
DC Voltage |  | 100 mV | 6.4 µV | /SJ-CAL-40417
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
DC Current |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
AC Voltage |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
AC Current |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
Analogue/digital multimeters | 40419 |  |  | Calibrator
DC Voltage |  | (±) |  | /SJ-CAL-40419
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
AC Voltage |  | (50 Hz ~ 10 kHz)
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Analogue/digital multimeters | 40419 |  |  | Calibrator
DC Current |  | (±) |  | /SJ-CAL-40419
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
AC Current |  | (50 Hz ~ 1 kHz)
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
(1 ~ 5) kHz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
Resistance |  | 1 Ω | 30 µΩ
|  | (1 ~ 10) Ω | 0.14 mΩ
|  | (10 ~ 100) Ω | 1.4 mΩ
|  | (0.1 ~ 1) kΩ | 14 mΩ
|  | (1 ~ 10) kΩ | 0.14 Ω
|  | (10 ~ 100) kΩ | 1.4 Ω
|  | (0.1 ~ 1) MΩ | 20 Ω
|  | (1 ~ 10) MΩ | 0.26 kΩ
|  | (10 ~ 100) MΩ | 2.8 kΩ
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Oscilloscopes | 40421 |  |  | Calibrator
DC Voltage |  | (±) |  | /SJ-CAL-40421
|  | 1 mV | 47 µV
|  | (1 ~ 10) mV | 51 µV
|  | (10 ~ 100) mV | 93 µV
|  | (0.1 ~ 1) V | 0.52 mV
|  | (1 ~ 10) V | 4.7 mV
|  | (10 ~ 100) V | 47 mV
|  | (100 ~ 120) V | 56 mV
AC Voltage(Square wave) |  | 1 kHz
|  | 2 mV | 49 µV
|  | (2 ~ 6) mV | 52 µV
|  | (6 ~ 12) mV | 58 µV
|  | (12 ~ 30) mV | 75 µV
|  | (30 ~ 60) mV | 0.11 mV
|  | (60 ~ 120) mV | 0.17 mV
|  | (120 ~ 300) mV | 0.34 mV
|  | (300 ~ 600) mV | 0.61 mV
|  | (0.6 ~ 1.2) V | 1.2 mV
|  | (1.2 ~ 3) V | 2.9 mV
|  | (3 ~ 6) V | 5.7 mV
|  | (6 ~ 12) V | 12 mV
|  | (12 ~ 30) V | 28 mV
|  | (30 ~ 60) V | 56 mV
|  | (60 ~ 120) V | 0.12 V
Bandwidth |  | 600 mV
|  | 50 kHz | 8.5 mV
|  | (50 ~ 100) kHz | 9.0 mV
|  | (0.1 ~ 10) MHz | 14 mV
|  | (10 ~ 60) MHz | 20 mV
|  | (60 ~ 600) MHz | 21 mV
|  | (600 ~ 1 000) MHz | 26 mV
|  | (1 000 ~ 1 100) MHz | 27 mV
Time mark |  | 1 ns | 0.62 ps
|  | (1 ~ 5) ns | 0.62 ps
|  | (5 ~ 50) ns | 6.2 ps
|  | (50 ~ 500) ns | 62 ps
|  | (0.5 ~ 5) µs | 0.62 ns
|  | (5 ~ 50) µs | 6.2 ns
|  | (50 ~ 500) µs | 62 ns
|  | (0.5 ~ 5) ms | 0.62 µs
|  | (5 ~ 50) ms | 6.2 µs
|  | (50 ~ 500) ms | 62 µs
|  | (0.5 ~ 5) s | 0.62 ms
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Voltage/current recorders | 40424 |  |  | Calibrator
DC Voltage |  | (±) |  | /SJ-CAL-40424
|  | 10 mV | 1.2 µV
|  | (10 ~ 100) mV | 2.1 µV
|  | (0.1 ~ 1) V | 8.8 µV
|  | (1 ~ 10) V | 84 µV
|  | (10 ~ 100) V | 1.2 mV
|  | (100 ~ 1 000) V | 12 mV
AC Voltage |  | (50 Hz ~ 10 kHz)
|  | 10 mV | 9.3 µV
|  | (10 ~ 100) mV | 21 µV
|  | (0.1 ~ 1) V | 0.15 mV
|  | (1 ~ 10) V | 1.4 mV
|  | (10 ~ 100) V | 14 mV
|  | (100 ~ 1 000) V | 0.23 V
DC Current |  | (±)
|  | 10 µA | 8.2 nA
|  | (10 ~ 100) µA | 19 nA
|  | (0.1 ~ 1) mA | 0.11 µA
|  | (1 ~ 10) mA | 1.1 µA
|  | (10 ~ 100) mA | 11 µA
|  | (0.1 ~ 1) A | 0.16 mA
|  | (1 ~ 10) A | 3.0 mA
|  | (10 ~ 20) A | 19 mA
|  | (20 ~ 30) A | 28 mA
AC Current |  | (50 Hz ~ 1 kHz)
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 53 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 21 µA
|  | (0.1 ~ 1) A | 0.29 mA
|  | (1 ~ 10) A | 3.3 mA
|  | (10 ~ 20) A | 22 mA
|  | (20 ~ 30) A | 29 mA
(1 ~ 5) kHz
|  | 10 µA | 34 nA
|  | (10 ~ 100) µA | 78 nA
|  | (0.1 ~ 1) mA | 0.35 µA
|  | (1 ~ 10) mA | 3.4 µA
|  | (10 ~ 100) mA | 32 µA
|  | (0.1 ~ 1) A | 0.32 mA
|  | (1 ~ 10) A | 4.4 mA
|  | (10 ~ 20) A | 0.10 A
|  | (20 ~ 30) A | 0.15 A
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Relay test sets | 40425 |  |  | Digital Multimeter
DC Voltage |  | 100 mV | 6.4 µV | /SJ-CAL-40425
|  | (0.1 ~ 1) V | 62 µV
|  | (1 ~ 10) V | 0.62 mV
|  | (10 ~ 100) V | 6.2 mV
|  | (100 ~ 1 000) V | 62 mV
DC Current |  | 100 mA | 14 µA
|  | (0.1 ~ 1) A | 0.24 mA
|  | (1 ~ 10) A | 3.1 mA
|  | (10 ~ 100) A | 60 mA
AC Voltage |  | (40 Hz ~ 1 kHz)
|  | 100 mV | 20 µV
|  | (0.1 ~ 1) V | 0.13 mV
|  | (1 ~ 10) V | 1.3 mV
|  | (10 ~ 100) V | 15 mV
|  | (100 ~ 1 000) V | 0.19 V
AC Current |  | (40 Hz ~ 1 kHz)
|  | 100 mA | 0.12 mA
|  | (0.1 ~ 1) A | 1.3 mA
|  | (1 ~ 10) A | 13 mA
|  | (10 ~ 100) A | 0.13 A
Frequency |  | 10 Hz | 0.62 mHz
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
Time |  | 1 s | 3.2 ms
|  | (1 ~ 10) s | 8.2 ms
|  | (10 ~ 100) s | 61 ms
LF signal generators | 40426 |  |  | Frequency Counter
Frequency |  | 10 Hz | 0.62 mHz | /SJ-CAL-40426
|  | (10 ~ 100) Hz | 6.2 mHz
|  | (0.1 ~ 1) kHz | 62 mHz
|  | (1 ~ 10) kHz | 0.62 Hz
|  | (10 ~ 100) kHz | 6.2 Hz
|  | (0.1 ~ 1) MHz | 62 Hz
Level Flatness |  | 1 V
|  | 10 Hz | 0.70 mV
|  | (10 ~ 100) Hz | 0.63 mV
|  | (0.1 ~ 1) kHz | 0.63 mV
|  | (1 ~ 10) kHz | 0.64 mV
|  | (10 ~ 100) kHz | 0.92 mV
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
LF signal generators | 40426 |  |  | Frequency Counter
Attenuation |  | 1 kHz |  | /SJ-CAL-40426
0 dBV |  | 1 V | 0.63 mV
-10 dBV |  | 316.2 mV | 91 µV
-20 dBV |  | 100 mV | 64 µV
501. Contact thermometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Temperature generators: oven, | 50101 |  |  | SPRT, Thermocouple
furnaces,isothermal liquid |  |  |  | /SJ-CAL-50101
bath, ice-points, dry-block
calibrators
Temperature controlled |  | -39.5 ℃ ~ 250 ℃ | 1.4 ℃
chembers/ovens
dry-block calibrators |  | -39.5 ℃ ~ 400 ℃ | 0.06 ℃
Temperature indicators/ | 50102 |  |  | SPRT, Thermocouple,
recorders/
controllers, calibrators
Temperature indicators/
recorders/controllers
sensor included |  | -39.5 ℃ ~ 250 ℃ | 0.04 ℃ | Decade Resistor,
|  | 250 ℃ ~ 1 100 ℃ | 0.9 ℃ | Calibrator
|  | 1 100 ℃ ~ 1 300 ℃ | 2.6 ℃ | /SJ-CAL-50102
sensor excepted |  | -39.5 ℃ ~ 1 300 ℃ | 0.1 ℃
Glass thermometers: | 50103 |  |  | SPRT
liquid-in-glass, Beckmann |  | -39.5 ℃ ~ 250 ℃ | 0.05 ℃ | /SJ-CAL-50103
Resistance thermometers: SPRT, | 50104 |  |  | SPRT
IPRT, thermistors. Etc |  | -39.5 ℃ ~ 250 ℃ | 0.03 ℃ | /SJ-CAL-50104
Thermal expansion thermometers: | 50105 |  |  | SPRT
bimetal, gas or liquid type |  | -39.5 ℃ ~ 250 ℃ | 0.30 ℃ | /SJ-CAL-50105
Thermocouples: noble metal, | 50106 |  |  | SPRT, Thermocouple
base metal, pure metal, |  |  |  | /SJ-CAL-50106
special type, etc.
base metal, noble metal |  | -39.5 ℃ ~ 250 ℃ | 0.3 ℃
|  | 250 ℃ ~ 1 100 ℃ | 1.3 ℃
|  | 1 100 ℃ ~ 1 300 ℃ | 2.7 ℃
503. Humidity
Measured Quantity Instrument or Gauge | Field code | Range | Measurement Uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Relative humidity hygrometers; | 50302 |  |  | Dew-point hygrometers
polimer thinfilm, hair, etc. |  |  |  | /SJ-CAL-50302
polimer thinfilm |  | (30 ~ 95) % R.H. | 2.2
(Relative humidity) |  | (-30 ~ 70) ℃ | 0.6
hair |  | (30 ~ 80) % R.H. | 4.2
(Relative humidity) |  | (-20 ~ 50) ℃ | 1.4
Temperature humidity recorders; | 50304 |  |  | Dew-point hygrometers
Hygrothermograph, etc |  | (30 ~ 80) % R.H. | 4.2 | /SJ-CAL-50304
(Relative humidity) |  | (-20 ~ 50) ℃ | 1.4
Humidity generators; | 50306 |  |  | Dew-point hygrometers
two-pressure, |  |  |  | /SJ-CAL-50306
two-temperature, flow mixing
humidity gererator, constant
temperature and humidity
chamber, etc.
constant temperature and |  | (30 ~ 60) % R.H. | 2.4
humidity chamber |  | (60 ~ 97) % R.H. | 3.4
(Relative humidity) |  | (-39.5 ~ 250) ℃ | 0.7
√("0." 〖"3" 4〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"2.51" 〗^2+("0.003 " 0×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"15" 〗^2+("0.003 0" ×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"0.1" 5〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"1." 8〗^2+("0.003 1" ×𝑙〖0)〗^2 ) " μm "
(𝑙0 :" mm")
√(〖"11.3" 〗^2+("0.004 " 0×𝑙〖0)〗^2 ) " μm "
(𝑙0 :" mm")
√(〖"0.77" 〗^2+("0.003 " 2×𝑙0)^2 ) "μm "
(𝑙0 :" mm")
√(〖"0.71" 〗^2+("0.04" 5×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"1.8" 〗^2+("0.003 0" ×𝑙〖0)〗^2 ) " μm "
(𝑙0 :" mm")
√(〖"1." 5〗^2+("0.003 0" ×𝑙0)^2 ) " μm "
(𝑙0 :" mm")
√(〖"0.92" 〗^2+("0.002 9" ×𝑙0)^2 ) "μm "
(𝑙0 :" mm")
√(〖"1." 9〗^2+("0.003 0" ×𝑙0)^2 ) "μm "
(𝑙0 :" mm")